Parada, A SPI Physics Prep Workbook

Modules built to prepare you for test taking strategies, tricky questions, and confusing topics.

Purpose: Build a foundation for solving SPI-style questions that test physical relationships, assumptions, calculations, and distractor recognition.

How to use this module

- Review the concept summaries and formulas.

- Answer each question before looking at the key.

- Record why you selected your answer, not just the letter.

- Review the distractor analysis, especially when you guessed correctly.

- Return to missed questions after a delay and solve them again without notes.

Part A — Essential concepts

1. Frequency, period, wavelength and propagation speed

- Frequency (f): Number of cycles per second. Unit: hertz (Hz).

- Period (T): Time required for one cycle. Unit: seconds.

- Wavelength (λ): Physical distance occupied by one cycle. Unit: meters or millimeters.

- Propagation speed (c): Speed at which sound travels through a medium. Unit: m/s.

Key relationships:

f = 1/T

λ = c/f

c = fλ

For diagnostic ultrasound, the commonly used average soft-tissue propagation speed is 1,540 m/s, or 1.54 mm/µs.

Exam trap: Propagation speed is primarily determined by the medium's properties, not by the transducer frequency. If frequency increases in the same medium, wavelength decreases.

2. What happens when frequency changes?

Assuming propagation speed remains constant:

| Parameter | If frequency doubles |
| --- | --- |
| Period | Halves |
| Wavelength | Halves |
| Propagation speed | Remains constant |
| Attenuation | Generally increases for the same path in soft tissue |
| Axial resolution | Can improve if spatial pulse length decreases |
| Penetration | Generally decreases |

Important qualification: A change in frequency alone does not guarantee a precise numerical change in resolution. The pulse characteristics and beam geometry matter.

3. Spatial pulse length and axial resolution

Spatial pulse length (SPL) is the physical length of a pulse in space.

SPL = number of cycles × wavelength

Axial resolution = SPL ÷ 2

To distinguish two reflectors located one behind the other, the returning echoes must be sufficiently separated.

Exam trap: Axial resolution is not the same as lateral resolution. Axial resolution depends on spatial pulse length. Lateral resolution depends mainly on beam width.

4. Pulsed-wave timing

- Pulse duration (PD): How long one pulse lasts.

- Pulse repetition period (PRP): Time from the start of one pulse to the start of the next.

- Pulse repetition frequency (PRF): Number of pulses transmitted per second.

- Duty factor (DF): Fraction of time the system is transmitting.

PRF = 1/PRP

DF = PD/PRP

For pulse-echo imaging, the system must allow time for echoes from the maximum imaging depth to return before the next pulse is transmitted.

PRP ≈ 2 × depth / propagation speed

For a propagation speed of 1.54 mm/µs:

PRP (µs) ≈ depth (mm) × 2 / 1.54

Therefore, greater imaging depth requires a longer listening period and limits the maximum usable PRF.

Exam trap: PRF and PRP move in opposite directions. If PRF increases, PRP decreases.

Part B — Exam-style multiple-choice questions

Choose the single best answer. Unless otherwise stated, assume conventional pulse-echo imaging and average soft-tissue propagation speed.

Question 1 — Direct relationship

A sound wave travels through the same tissue at a constant propagation speed. Its frequency increases from 4 MHz to 8 MHz. What happens to its wavelength?

A. It doubles.
B. It is reduced by half.
C. It remains unchanged.
D. It increases fourfold.

Question 2 — Identify the incorrect statement

Which statement about diagnostic ultrasound is incorrect?

A. Frequency is measured in hertz.
B. Period and frequency are inversely related.
C. Propagation speed is primarily determined by the medium.
D. Increasing frequency in the same medium increases wavelength.

Question 3 — Calculation with units

A sound wave has a frequency of 5 MHz and travels through soft tissue at 1,540 m/s. What is its approximate wavelength?

A. 0.031 mm
B. 0.154 mm
C. 0.308 mm
D. 3.08 mm

Question 4 — Resolution misconception

Which change most directly improves axial resolution?

A. Increasing beam width.
B. Increasing spatial pulse length.
C. Decreasing spatial pulse length.
D. Increasing imaging depth.

Question 5 — Pulse timing

A system's PRF increases from 2,000 pulses/s to 4,000 pulses/s. What happens to PRP?

A. It doubles from 0.25 ms to 0.50 ms.
B. It halves from 0.50 ms to 0.25 ms.
C. It remains unchanged.
D. It doubles from 0.50 ms to 1.00 ms.

Question 6 — Depth and timing

A sonographer increases the maximum imaging depth while all other relevant system constraints remain unchanged. Which change is generally necessary to prevent echoes from the previous pulse from overlapping the next transmission?

A. Increase the maximum PRF.
B. Decrease the PRP.
C. Allow a longer listening period, reducing the maximum PRF.
D. Increase the frequency to maintain the same PRF.

Question 7 — Frequency trade-off

A sonographer switches to a higher-frequency transducer while imaging the same deep structure. Which outcome is most likely?

A. Greater penetration and poorer potential axial resolution.
B. Less attenuation and greater penetration.
C. Greater attenuation and potentially better axial resolution.
D. Increased propagation speed and unchanged wavelength.

Question 8 — Two-step reasoning

A pulse contains three cycles. The wavelength is 0.30 mm. What are the spatial pulse length and ideal axial resolution?

A. SPL = 0.10 mm; axial resolution = 0.05 mm.
B. SPL = 0.90 mm; axial resolution = 0.45 mm.
C. SPL = 0.90 mm; axial resolution = 0.90 mm.
D. SPL = 0.30 mm; axial resolution = 0.15 mm.

Question 9 — Duty factor

A system transmits for 2 µs during a pulse repetition period of 1,000 µs. What is the duty factor?

A. 0.002, or 0.2%.
B. 0.02, or 2%.
C. 0.2, or 20%.
D. 2, or 200%.

Question 10 — Multiple relationships

A sonographer decreases the operating frequency while maintaining the same propagation medium. Which combination is most consistent with the change?

A. Shorter wavelength, greater attenuation, shorter period.
B. Longer wavelength, lower attenuation, longer period.
C. Longer wavelength, greater attenuation, shorter period.
D. Unchanged wavelength, lower propagation speed, longer period.

Question 11 — A subtle assumption

Two transducers operate in the same tissue at 3 MHz and 6 MHz. Which statement must be true under the same propagation-speed assumption?

A. The 6 MHz wave travels twice as fast.
B. The 6 MHz wave has half the wavelength.
C. The 6 MHz wave has twice the period.
D. The 3 MHz wave has half the wavelength.

Question 12 — Best answer versus absolute rule

A student states, “Higher frequency always gives better resolution.” Which response is the most scientifically accurate?

A. Correct, because wavelength determines every type of resolution.
B. Incorrect, because frequency has no relationship to resolution.
C. Incomplete, because higher frequency can improve axial resolution, but lateral and elevational resolution depend on beam geometry.
D. Correct, provided the imaging depth is increased.

Part C — Answer key and distractor analysis

1. Correct answer: B

Wavelength equals propagation speed divided by frequency. Doubling frequency while keeping propagation speed constant halves wavelength.

- A reverses the relationship.

- C ignores the inverse relationship.

- D incorrectly applies a fourfold change.

Test-taking strategy: When one variable is in the denominator, increasing that variable reduces the result if the numerator remains constant.

2. Correct answer: D

Increasing frequency reduces wavelength when propagation speed remains constant.

A, B and C are correct statements. This is a classic exception question: identify the one statement that contradicts the physics.

Test-taking strategy: For “incorrect” questions, evaluate each option independently rather than selecting the first familiar statement.

3. Correct answer: C

Use λ = c/f.

Convert frequency: 5 MHz = 5,000,000 Hz.

λ = 1,540 / 5,000,000 m = 0.000308 m = 0.308 mm.

Common trap: Forgetting to convert meters to millimeters or mixing MHz with Hz.

4. Correct answer: C

Axial resolution equals spatial pulse length divided by two. Shortening the pulse allows two reflectors along the beam axis to be distinguished when closer together.

A concerns lateral resolution. B worsens axial resolution. D does not directly improve axial resolution.

5. Correct answer: B

PRP = 1/PRF.

At 2,000 pulses/s, PRP = 0.0005 s = 0.50 ms.

At 4,000 pulses/s, PRP = 0.00025 s = 0.25 ms.

Doubling PRF halves PRP.

Common trap: Confusing a frequency with a time interval.

6. Correct answer: C

Greater depth requires more time for the pulse to travel to the reflector and return. A longer listening interval is therefore needed, limiting the maximum PRF.

A and B would make pulse timing less suitable for greater depths. D confuses frequency with propagation time.

7. Correct answer: C

Higher frequency generally produces greater attenuation in soft tissue and can improve axial resolution through a shorter wavelength and shorter pulse.

The trade-off is reduced penetration. Propagation speed does not double merely because transducer frequency doubles.

8. Correct answer: B

SPL = cycles × wavelength = 3 × 0.30 mm = 0.90 mm.

Axial resolution = SPL/2 = 0.45 mm.

Common trap: Reporting the spatial pulse length as the axial resolution without dividing by two.

9. Correct answer: A

DF = PD/PRP = 2/1,000 = 0.002.

As a percentage, 0.002 × 100 = 0.2%.

Common trap: Confusing the decimal ratio with the percentage, or assuming duty factor must be greater than 1.

10. Correct answer: B

Reducing frequency increases wavelength and increases period because period and frequency are inversely related. Attenuation generally decreases at lower frequencies in the same tissue.

This question combines three relationships, so students must evaluate the full answer rather than one correct phrase.

11. Correct answer: B

With constant propagation speed, doubling frequency halves wavelength.

The wave's propagation speed is governed primarily by the medium, and period decreases rather than increases.

12. Correct answer: C

Frequency and wavelength are important to axial resolution, but they do not independently determine every type of resolution. Lateral resolution is related to beam width, while elevational resolution depends on slice thickness.

Exam lesson: Be cautious with absolute words such as always, never, and only. They are not automatically wrong, but they require careful verification.

Part D — Cross-concept review

The frequency trade-off

When frequency increases, students should connect the following concepts:

- Wavelength decreases.

- Period decreases.

- Attenuation generally increases in soft tissue.

- Penetration generally decreases.

- Axial resolution can improve when pulse length decreases.

- Lateral resolution depends on beam width and focus, not frequency alone.

The depth–PRF trade-off

When maximum imaging depth increases:

- The pulse must travel farther.

- Echo-return time increases.

- The system must allow a longer listening interval.

- The maximum usable PRF decreases.

- Lower PRF can also limit the maximum unambiguous Doppler shift, all else being equal.

The pulse-length relationship

Shorter SPL improves axial resolution because:

Axial resolution = SPL/2

SPL = number of cycles × wavelength

Therefore, axial resolution can be improved by reducing wavelength, reducing the number of cycles per pulse, or both. Damping is important because it helps produce shorter pulses.

Part E — Final self-check

Before moving on, a student should be able to explain, without looking at the answer key:

- Why frequency and period are inversely related.

- Why wavelength changes when frequency changes in the same medium.

- Why higher frequency creates a penetration trade-off.

- Why axial and lateral resolution cannot be treated as interchangeable.

- Why increasing depth limits PRF.

- How to convert MHz to Hz and meters to millimeters.

- Why a numerically correct formula can still produce an incorrect multiple-choice answer if the units are wrong.

Recommended mastery standard: Aim for at least 10 of 12 correct initially, then revisit every missed or guessed question and explain the reasoning aloud before proceeding.